## Summary
Sample Location: |  | Various offshore/inshore sediment sites located inside Sinclair Inlet along Pier 7 [Operable Unit B (OU-B)] :
Naval Shipyard (NSY) Puget Sound, Bremerton, Washington (USA)
Sampling Date: 28 Oct. 2010
Sample Type: Sediment (wet), grab, core
Analysis: Poly-Chlorinated Biphenyls (PCBs) by Enzyme-Linked Immuno Sorbent Assay (ELISA) Methods
ImmunoAssay Application: EPA Methods: 4020 (PCB) For Soils (Modified for Marine Sediments)
Analysis Date: 31 Oct. 2010
Analyst: Joel M. Guerrero, Scientist (Environmental Chemistry), U.S. Navy Space & Naval Warfare Systems Center (SPAWARSYSCEN) Pacific,
Advanced Systems & Applied Sciences Division, Applied Systems & Environmental Sciences Branch (Code 71751), (619) 553-4169
Analysis Location: Building 147, PSNS&IMF Foundry Bremerton (WA)
POCs: Victoria Kirtay, Scientist (Marine Geochemistry), SPAWARSYSCEN, Pacific [SSC-Pacific], Code 71751: (619) 553-1395
Joel M. Guerrero, Scientist (Environmental Chemistry), SPAWARSYSCEN Pacific [SSC-Pacific], Code 71751: (619) 553-4169
Method Summary: The ImmunoAssay (IA) test kit used for this method are commercially available (RaPIDTM Assay, Stategic Diagnostics Inc.,
www.sdix.com). In general, the IA method is perfomed using a sample extract. The sample and an enzyme conjugate are added
to an antibody. The enzyme conjugate "competes" with the contaminant of interest (PCB) present in the sample
for binding to the antibody. The assay test is interpreted by comparing the response produced by the sample test
to the response produced by testing a range of kit-supplied standards simultaneously.
Sample Preparation Procedures:
(1) Pre-processing: mix wet sediment thoroughly with clean sampler spoons. using filter paper (grade 13V) and heavy-duty lab towels, about 20-30 grams of
wet sample are dewatered and/or air-dried to reduce sediment moisture content (at least less than 30%). observe and record other sediment
meta data (i.e. color, grain size, smell, oil sheen, debris, shell hash, etc.)
(2) Sub-sampling: weigh and record approx. 10 grams of "dry" (at least 70% or more) sediment using top loading balance. transfer into
50 ml centrifuge tubes (pre-cleaned/sterile).
(3) Extraction: add 20 ml of methanol (MeOH, Aldrich 99.9+% capillary GC grade). cap centrifuge tube tight. vortex slurry for 1-2 seconds
to loosen particles and break-up large sediment clumps. shake the mixture vigorously and continously for at least 1 hour in a rotator/shaker
set at maximum amplitude (400 rpm).
(4) Separation: check the tube cap again for a tight seal. centrifuge the slurry for at least 10-15 minutes (EIC/Sorvall) at 3000 rpm.
transfer/pour MeOH extract into a syringeless filter vial and filter through (GFC/Whatman, 0.45 um pore size) into amber, teflon-lined vial.
record extract color. cap and seal the vial, keep dark and cool (2-8 degrees C).
(5) Dilution: pipet (using eppendorf) 50 μL of MeOH extract into glass test tube (pre-cleaned with screw cap) containing 5 mL of sample
diluent (reagent buffer saline solution having preservatives and stabilizers without any detectable PCBs).
cap test tubes and vortex to mix.
(6) Assay: pre-label and prepare polystyrene test tubes (allow for field/sample/extract duplicates for each assay run).
perform immunoassay (IA) procedures by following SDI standard soil methods for PCBs.
ImmunoAssay (IA) ELISA Protocols: The sample (with the unknown contaminant concentration; ex. PCB) is analyzed by the addition of an
enzyme conjugate (labeled PCB).This is followed by addition of paramagnetic particles with anti-bodies specific to "both" PCBs.
In relatively proportional concentrations, both the sample PCBs and the "labeled" PCBs (conjugate) compete for the binding sites on the magnetic particles.
After an incubation period, a magnetic field is applied to hold (in-place) the magnetic particles having the sample PCB and its "labeled" PCB analog to bind
with the antibodies. Any unbound reagents are decanted and washed repeatedly. PCBs in the mixture are detected with the addition of an enzyme substrate
(color solution) containing a chromagen which specifically react to the "labelled" PCBs. After another incubation, the reaction is stopped
and stabilized by addition of acid (stopping solution). Since the labelled PCBs and sample PCBs are in competition (proportionally)
with the binding sites, the color developed at the end of reaction is inversely proportional to the concentration of PCBs in the sample.
This color response is measured by a spectrophotometer (set at 450 nm) and compared to the responses taken from
a calibrated series of known pcb standards (kit-supplied) to determine the equivalent PCB (as Aroclor 1254) concentration of the sample.
Calculations: The dry-weight corrected organic contaminant (OC) concentration in sediment is calculated (according to eq. 1 and eq. 2)
by multiplying the IA result by factors that are introduced by the sample pre-processing procedures (collection, extraction, extract dilution steps above).
| OCconc = IA x (VMeOH / Wdrysed) x df |  | (eq.1)
|  | df = (Vext + Vdil) / Vext | (eq.2)
where;
| OCconc: | PCB sediment concentration (mg/kg)
| IA: | ImmunoAssay result (mg/L)
| VMeOH: | Volume of methanol (extractant)
| Wdrysed: | Weight of dry sediment
| df: | Dilution factor
| Vext: | Volume of dilution extract
| Vdil: | Volume of diluent
Colorimetry: Absorbance values at l = 450 nm were measured/recorded using a field-portable/lab benchtop HACH Model DR/2010
direct-reading spectrophotometer (HACH Co., Loveland, CO). This analyzer is a microprocessor-controlled, single-beam instrument
with a Littrow Prism monochromator design which uses an ultra-violet (UV) enhanced silicon photodiode for detection.
It is manually programmed to display results in raw absorbances. The HACH DR/2010 is calibrated at the start of each assay run.
It is zeroed for absorbance measurements and a series of assay method calibrations are performed using a contaminant-free 0 (zero) standard,
a range of 3 standard concentrations, and a control. Method/equipment blanks and marine sediment reference materials
[NRCC HS-1(PCB)] are also used for standardization. Measurements of raw absorbance data are serially down-loaded
to a laptop computer using a WindowsTM application HACHLinkTM data acquisition software. Serial data, collected as text files, are transferred
and copied into a spreadsheet program (ExcelTM) for reduction. A calibration curve (with R2= 0.9) is generated for each assay run

## ImmunoAssay (IA) Results
Field ID | tPCB
SAMPLE LABEL | (mg/Kg) | Q | Stdev | %RSD
P7-T1-1 | 234.3457
P7-T1-2 | 191.9109
P7-T1-3 | 90.8979
P7-T1-4 | 45.6964 | U
P7-T1-5 | 151.8774 |  | 46.6869 | 30.74%
P7-T2-1 | 176.2019
P7-T2-2 | 150.7784
P7-T2-3 | 27.7236 | U
P7-T2-4 | 260.9067
P7-T2-5 | 58.049 | U
P7-T3-1 | 169.7766
P7-T3-2 | 96.2914 | J
P7-T3-3 | 284.7996
P7-T3-4 | 74.2266 | U
P7-T3-5 | 48.402 |  | 24.9813 | 51.61%
P7-T4-1 | 340.6951
P7-T4-2 | 140.4015 | J
P7-T4-3 | 113.1814 | J
P7-T4-4 | 133.1015 | J
P7-T4-5 | 10.9711 | U
P7-T5-1 | 111.0749 | J
P7-T5-2 | 594.0079 |  | 76.428 | 12.87%
P7-T5-3 | 159.287
P7-T5-4 | 150.0182
P7-T5-5 | 73.7231 | U
P7-T6-1 | 243.2908
P7-T6-2 | 261.6431
P7-T6-3 | 6650.4322 |  | 1423.8799 | 21.41%
P7-T6-4 | 305.3487
P7-T6-5 | 438.6649
P7-T7-1 | 224.2514
P7-T7-2 | 261.4078
P7-T7-3 | 128.9571 | J
P7-T7-4 | 654.6141
P7-T7-5 | 192.8252
P7-T7-6 | 104.9565 | J
P7-T8-1 | 226.743
P7-T8-2 | 129.253 | J
P7-T8-3 | 162.6665
Definitions:
Stdev: Standard Deviation from duplicate assay analyses (n=2)
% RSD: Percent Relative Standard Deviation whereby; [{stdev/mean} * 100]
Q: Data Qualifiers: U = None-Detect, J = Estimated, E = Outside Linear Range, Blank = Detect
Label: P7 = Pier 7; Tx-x = Transect № (1…10) - Sample № (1…6)
Field ID | tPCB
SAMPLE LABEL | (μg/Kg) | Q | Stdev | %RSD
P7-T9-1 | 91.4489 | J | 53.2678 | 58.25%
P7-T9-2 | 73.888 | U
P7-T9-3 | 74.3209 | U
P7-T9-4 | 73.3692 | U
P7-T9-5 | 84.3271 | J
P7-T9-6 | 80.1227 | J
P7-T10-1 | 160.8839
P7-T10-2 | 23.5901 | U
P7-T10-3 | 126.4404 | J
P7-T10-4 | 91.5388 | J
P7-T10-5 | 134.3635 | J
P7-T10-6 | 115.4552 | J | 8.9471 | 7.749%
Definitions:
Stdev: Standard Deviation from duplicate assay analyses (n=2)
% RSD: Percent Relative Standard Deviation whereby; [{stdev/mean} * 100]
Q: Data Qualifiers: U = None-Detect, J = Estimated, E = Outside Linear Range, Blank = Detect
Label: P7 = Pier 7; Tx-x = Transect № (1…10) - Sample № (1…6)

## Assay Duplicates
SAMPLE |  | tPCB (mg/Kg)
LABEL | 1 | 2 | 3 | Mean | Stdev | %RSD
P7-T1-1 | 234.3457 |  |  | 234.3457
P7-T1-2 | 191.9109 |  |  | 191.9109
P7-T1-3 | 90.8979 |  |  | 90.8979
P7-T1-4 | 45.6964 |  |  | 45.6964
P7-T1-5 | 118.8648 | 184.89 |  | 151.8774 | 46.6869 | 30.74%
P7-T2-1 | 176.2019 |  |  | 176.2019
P7-T2-2 | 150.7784 |  |  | 150.7784
P7-T2-3 | 27.7236 |  |  | 27.7236
P7-T2-4 | 260.9067 |  |  | 260.9067
P7-T2-5 | 58.049 |  |  | 58.049
P7-T3-1 | 169.7766 |  |  | 169.7766
P7-T3-2 | 96.2914 |  |  | 96.2914
P7-T3-3 | 284.7996 |  |  | 284.7996
P7-T3-4 | 74.2266 |  |  | 74.2266
P7-T3-5 | 30.7375 | 66.0665 |  | 48.402 | 24.9813 | 51.61%
P7-T4-1 | 340.6951 |  |  | 340.6951
P7-T4-2 | 140.4015 |  |  | 140.4015
P7-T4-3 | 113.1814 |  |  | 113.1814
P7-T4-4 | 133.1015 |  |  | 133.1015
P7-T4-5 | 10.9711 |  |  | 10.9711
P7-T5-1 | 111.0749 |  |  | 111.0749
P7-T5-2 | 539.9652 | 648.0507 |  | 594.0079 | 76.428 | 12.87%
P7-T5-3 | 159.287 |  |  | 159.287
P7-T5-4 | 150.0182 |  |  | 150.0182
P7-T5-5 | 73.7231 |  |  | 73.7231
P7-T6-1 | 243.2908 |  |  | 243.2908
P7-T6-2 | 261.6431 |  |  | 261.6431
P7-T6-3 | 5643.5971 | 7657.2673 |  | 6650.4322 | 1423.8799 | 21.41%
P7-T6-4 | 305.3487 |  |  | 305.3487
P7-T6-5 | 438.6649 |  |  | 438.6649
P7-T7-1 | 224.2514 |  |  | 224.2514
P7-T7-2 | 261.4078 |  |  | 261.4078
P7-T7-3 | 128.9571 |  |  | 128.9571
P7-T7-4 | 654.6141 |  |  | 654.6141
P7-T7-5 | 192.8252 |  |  | 192.8252
P7-T7-6 | 104.9565 |  |  | 104.9565
P7-T8-1 | 226.743 |  |  | 226.743
P7-T8-2 | 129.253 |  |  | 129.253
P7-T8-3 | 162.6665 |  |  | 162.6665
Definitions:
Mean: Arithmethic Mean
Stdev: Standard Deviation from duplicate assay analyses (n=2)
% RSD: Percent Relative Standard Deviation whereby; [{stdev/mean} * 100]
Label: P7 = Pier 7; Tx-x = Transect № (1…10) - Sample № (1…6)
SAMPLE |  | tPCB (mg/Kg)
LABEL | 1 | 2 | 3 | Mean | Stdev | %RSD
P7-T9-1 | 53.7829 | 129.115 |  | 91.4489 | 53.2678 | 58.25%
P7-T9-2 | 73.888 |  |  | 73.888
P7-T9-3 | 74.3209 |  |  | 74.3209
P7-T9-4 | 73.3692 |  |  | 73.3692
P7-T9-5 | 84.3271 |  |  | 84.3271
P7-T9-6 | 80.1227 |  |  | 80.1227
P7-T10-1 | 160.8839 |  |  | 160.8839
P7-T10-2 | 23.5901 |  |  | 23.5901
P7-T10-3 | 126.4404 |  |  | 126.4404
P7-T10-4 | 91.5388 |  |  | 91.5388
P7-T10-5 | 134.3635 |  |  | 134.3635
P7-T10-6 | 109.1287 | 121.7818 |  | 115.4552 | 8.9471 | 7.749%
Definitions:
Mean: Arithmethic Mean
Stdev: Standard Deviation from duplicate assay analyses (n=2)
% RSD: Percent Relative Standard Deviation whereby; [{stdev/mean} * 100]
Label: P7 = Pier 7; Tx-x = Transect № (1…10) - Sample № (1…6)

## PCB Assay Raw Absorbance Data
Assay Date | Time | abs | %T | SSC-Pacific LAB ID | Operator | HACH Model Serial № | Progam № Wavelength λ
2010-10-31 | 0:00 | 0 | 100 | zero | Joel Guerrero | DR/2010 SN:000200017081 | P0; 450nm; PCB
2010-10-31 | 0:00 | 1.166 | 6.83 | 0.0 std | Joel Guerrero | DR/2010 SN:000200017081 | P0; 450nm; PCB
2010-10-31 | 0:00 | 0.902 | 12.5 | 0.25 std | Joel Guerrero | DR/2010 SN:000200017081 | P0; 450nm; PCB
2010-10-31 | 0:00 | 0.664 | 21.7 | 1.00 std | Joel Guerrero | DR/2010 SN:000200017081 | P0; 450nm; PCB
2010-10-31 | 0:00 | 0.467 | 34.1 | 5.00 std | Joel Guerrero | DR/2010 SN:000200017081 | P0; 450nm; PCB
2010-10-31 | 0:00 | 0.521 | 30.1 | control | Joel Guerrero | DR/2010 SN:000200017081 | P0; 450nm; PCB
2010-10-31 | 0:00 | 1.075 | 8.42 | RB | Joel Guerrero | DR/2010 SN:000200017081 | P0; 450nm; PCB
2010-10-31 | 0:00 | 0.78 | 16.6 | P7-T1-1 | Joel Guerrero | DR/2010 SN:000200017081 | P0; 450nm; PCB
2010-10-31 | 0:00 | 0.707 | 19.6 | P7-T1-2 | Joel Guerrero | DR/2010 SN:000200017081 | P0; 450nm; PCB
2010-10-31 | 0:00 | 0.814 | 15.4 | P7-T1-3 | Joel Guerrero | DR/2010 SN:000200017081 | P0; 450nm; PCB
2010-10-31 | 0:00 | 0.898 | 12.6 | P7-T1-4 | Joel Guerrero | DR/2010 SN:000200017081 | P0; 450nm; PCB
2010-10-31 | 0:00 | 0.777 | 16.7 | P7-T1-5 | Joel Guerrero | DR/2010 SN:000200017081 | P0; 450nm; PCB
2010-10-31 | 0:00 | 0.72 | 19 | P7-T2-1 | Joel Guerrero | DR/2010 SN:000200017081 | P0; 450nm; PCB
2010-10-31 | 0:00 | 0.743 | 18.1 | P7-T2-2 | Joel Guerrero | DR/2010 SN:000200017081 | P0; 450nm; PCB
2010-10-31 | 0:00 | 0.95 | 11.2 | P7-T2-3 | Joel Guerrero | DR/2010 SN:000200017081 | P0; 450nm; PCB
2010-10-31 | 0:00 | 0.661 | 21.8 | P7-T2-4 | Joel Guerrero | DR/2010 SN:000200017081 | P0; 450nm; PCB
2010-10-31 | 0:00 | 0.871 | 13.4 | P7-T2-5 | Joel Guerrero | DR/2010 SN:000200017081 | P0; 450nm; PCB
2010-10-31 | 0:00 | 0.726 | 18.8 | P7-T3-1 | Joel Guerrero | DR/2010 SN:000200017081 | P0; 450nm; PCB
2010-10-31 | 0:00 | 0.806 | 15.6 | P7-T3-2 | Joel Guerrero | DR/2010 SN:000200017081 | P0; 450nm; PCB
2010-10-31 | 0:00 | 0.647 | 22.5 | P7-T3-3 | Joel Guerrero | DR/2010 SN:000200017081 | P0; 450nm; PCB
2010-10-31 | 0:00 | 0.84 | 14.4 | P7-T3-4 | Joel Guerrero | DR/2010 SN:000200017081 | P0; 450nm; PCB
2010-10-31 | 0:00 | 0.94 | 11.5 | P7-T3-5 | Joel Guerrero | DR/2010 SN:000200017081 | P0; 450nm; PCB
2010-10-31 | 0:00 | 0.619 | 24 | P7-T4-1 | Joel Guerrero | DR/2010 SN:000200017081 | P0; 450nm; PCB
2010-10-31 | 0:00 | 0.754 | 17.6 | P7-T4-2 | Joel Guerrero | DR/2010 SN:000200017081 | P0; 450nm; PCB
2010-10-31 | 0:00 | 0.784 | 16.4 | P7-T4-3 | Joel Guerrero | DR/2010 SN:000200017081 | P0; 450nm; PCB
2010-10-31 | 0:00 | 0.762 | 17.3 | P7-T4-4 | Joel Guerrero | DR/2010 SN:000200017081 | P0; 450nm; PCB
2010-10-31 | 0:00 | 1.025 | 9.45 | P7-T4-5 | Joel Guerrero | DR/2010 SN:000200017081 | P0; 450nm; PCB
2010-10-31 | 0:00 | 1.06 | 8.7 | 0.00 std | Joel Guerrero | DR/2010 SN:000200017081 | P0; 450nm; PCB
2010-10-31 | 0:00 | 0.866 | 13.6 | 0.25 std | Joel Guerrero | DR/2010 SN:000200017081 | P0; 450nm; PCB
2010-10-31 | 0:00 | 0.643 | 22.8 | 1.00 std | Joel Guerrero | DR/2010 SN:000200017081 | P0; 450nm; PCB
2010-10-31 | 0:00 | 0.451 | 35.4 | 5.00 std | Joel Guerrero | DR/2010 SN:000200017081 | P0; 450nm; PCB
2010-10-31 | 0:00 | 0.509 | 31 | control | Joel Guerrero | DR/2010 SN:000200017081 | P0; 450nm; PCB
2010-10-31 | 0:00 | 1.034 | 9.25 | MB | Joel Guerrero | DR/2010 SN:000200017081 | P0; 450nm; PCB
2010-10-31 | 0:00 | 0.763 | 17.3 | P7-T5-1 | Joel Guerrero | DR/2010 SN:000200017081 | P0; 450nm; PCB
2010-10-31 | 0:00 | 0.529 | 29.6 | P7-T5-2 | Joel Guerrero | DR/2010 SN:000200017081 | P0; 450nm; PCB
2010-10-31 | 0:00 | 0.715 | 19.3 | P7-T5-3 | Joel Guerrero | DR/2010 SN:000200017081 | P0; 450nm; PCB
2010-10-31 | 0:00 | 0.723 | 18.9 | P7-T5-4 | Joel Guerrero | DR/2010 SN:000200017081 | P0; 450nm; PCB
2010-10-31 | 0:00 | 0.812 | 15.4 | P7-T5-5 | Joel Guerrero | DR/2010 SN:000200017081 | P0; 450nm; PCB
2010-10-31 | 0:00 | 0.653 | 22.2 | P7-T6-1 | Joel Guerrero | DR/2010 SN:000200017081 | P0; 450nm; PCB
2010-10-31 | 0:00 | 0.643 | 22.7 | P7-T6-2 | Joel Guerrero | DR/2010 SN:000200017081 | P0; 450nm; PCB
2010-10-31 | 0:00 | 0.21 | 61.6 | P7-T6-3 | Joel Guerrero | DR/2010 SN:000200017081 | P0; 450nm; PCB
2010-10-31 | 0:00 | 0.619 | 24 | P7-T6-4 | Joel Guerrero | DR/2010 SN:000200017081 | P0; 450nm; PCB
2010-10-31 | 0:00 | 0.562 | 27.4 | P7-T6-5 | Joel Guerrero | DR/2010 SN:000200017081 | P0; 450nm; PCB
2010-10-31 | 0:00 | 0.666 | 21.6 | P7-T7-1 | Joel Guerrero | DR/2010 SN:000200017081 | P0; 450nm; PCB
2010-10-31 | 0:00 | 0.643 | 22.7 | P7-T7-2 | Joel Guerrero | DR/2010 SN:000200017081 | P0; 450nm; PCB
2010-10-31 | 0:00 | 0.743 | 18.1 | P7-T7-3 | Joel Guerrero | DR/2010 SN:000200017081 | P0; 450nm; PCB
2010-10-31 | 0:00 | 0.5 | 31.6 | P7-T7-4 | Joel Guerrero | DR/2010 SN:000200017082 | P0; 450nm; PCB
2010-10-31 | 0:00 | 0.688 | 20.5 | P7-T7-5 | Joel Guerrero | DR/2010 SN:000200017083 | P0; 450nm; PCB
2010-10-31 | 0:00 | 0.769 | 17 | P7-T7-6 | Joel Guerrero | DR/2010 SN:000200017084 | P0; 450nm; PCB
2010-10-31 | 0:00 | 0.664 | 21.7 | P7-T8-1 | Joel Guerrero | DR/2010 SN:000200017085 | P0; 450nm; PCB
2010-10-31 | 0:00 | 0.743 | 18.1 | P7-T8-2 | Joel Guerrero | DR/2010 SN:000200017086 | P0; 450nm; PCB
2010-10-31 | 0:00 | 0.711 | 19.4 | P7-T8-3 | Joel Guerrero | DR/2010 SN:000200017087 | P0; 450nm; PCB
2010-10-31 | 0:00 | 0.571 | 26.9 | HS1 | Joel Guerrero | DR/2010 SN:000200017088 | P0; 450nm; PCB
Assay Date | Time | abs | %T | SSC-Pacific LAB ID | Operator | HACH Model Serial № | Progam № Wavelength λ
2010-10-31 | 0:00 | 0 | 100 | zero | Joel Guerrero | DR/2010 SN:000200016870 | P0; 450nm; PCB
2010-10-31 | 0:00 | 1.438 | 3.65 | 0.00 std | Joel Guerrero | DR/2010 SN:000200016870 | P0; 450nm; PCB
2010-10-31 | 0:00 | 1.096 | 8.03 | 0.25 std | Joel Guerrero | DR/2010 SN:000200016870 | P0; 450nm; PCB
2010-10-31 | 0:00 | 0.784 | 16.5 | 1.00 std | Joel Guerrero | DR/2010 SN:000200016870 | P0; 450nm; PCB
2010-10-31 | 0:00 | 0.531 | 29.5 | 5.00 std | Joel Guerrero | DR/2010 SN:000200016870 | P0; 450nm; PCB
2010-10-31 | 0:00 | 0.602 | 25 | control | Joel Guerrero | DR/2010 SN:000200016870 | P0; 450nm; PCB
2010-10-31 | 0:00 | 1.304 | 4.96 | RB | Joel Guerrero | DR/2010 SN:000200016870 | P0; 450nm; PCB
2010-10-31 | 0:00 | 1.067 | 8.56 | P7-T9-1 | Joel Guerrero | DR/2010 SN:000200016870 | P0; 450nm; PCB
2010-10-31 | 0:00 | 1.016 | 9.65 | P7-T9-2 | Joel Guerrero | DR/2010 SN:000200016870 | P0; 450nm; PCB
2010-10-31 | 0:00 | 1.015 | 9.66 | P7-T9-3 | Joel Guerrero | DR/2010 SN:000200016870 | P0; 450nm; PCB
2010-10-31 | 0:00 | 1.017 | 9.61 | P7-T9-4 | Joel Guerrero | DR/2010 SN:000200016870 | P0; 450nm; PCB
2010-10-31 | 0:00 | 0.993 | 10.2 | P7-T9-5 | Joel Guerrero | DR/2010 SN:000200016870 | P0; 450nm; PCB
2010-10-31 | 0:00 | 1.001 | 9.99 | P7-T9-6 | Joel Guerrero | DR/2010 SN:000200016870 | P0; 450nm; PCB
2010-10-31 | 0:00 | 0.874 | 13.4 | P7-T10-1 | Joel Guerrero | DR/2010 SN:000200016870 | P0; 450nm; PCB
2010-10-31 | 0:00 | 1.18 | 6.61 | P7-T10-2 | Joel Guerrero | DR/2010 SN:000200016870 | P0; 450nm; PCB
2010-10-31 | 0:00 | 0.92 | 12 | P7-T10-3 | Joel Guerrero | DR/2010 SN:000200016870 | P0; 450nm; PCB
2010-10-31 | 0:00 | 0.978 | 10.5 | P7-T10-4 | Joel Guerrero | DR/2010 SN:000200016870 | P0; 450nm; PCB
2010-10-31 | 0:00 | 0.908 | 12.4 | P7-T10-5 | Joel Guerrero | DR/2010 SN:000200016870 | P0; 450nm; PCB
2010-10-31 | 0:00 | 0.947 | 11.3 | P7-T10-6 | Joel Guerrero | DR/2010 SN:000200016870 | P0; 450nm; PCB
2010-10-31 | 0:00 | 1.297 | 5.05 | 0.00 std | Joel Guerrero | DR/2010 SN:000200016870 | P0; 450nm; PCB
2010-10-31 | 0:00 | 1.043 | 9.06 | 0.25 std | Joel Guerrero | DR/2010 SN:000200016870 | P0; 450nm; PCB
2010-10-31 | 0:00 | 0.833 | 14.7 | 1.00 std | Joel Guerrero | DR/2010 SN:000200016870 | P0; 450nm; PCB
2010-10-31 | 0:00 | 0.546 | 28.5 | 5.00 std | Joel Guerrero | DR/2010 SN:000200016870 | P0; 450nm; PCB
2010-10-31 | 0:00 | 0.621 | 24 | control | Joel Guerrero | DR/2010 SN:000200016870 | P0; 450nm; PCB
2010-10-31 | 0:00 | 1.288 | 5.07 | MB | Joel Guerrero | DR/2010 SN:000200016870 | P0; 450nm; PCB
2010-10-31 | 0:00 | 0.75 | 17.8 | PQ1 | Joel Guerrero | DR/2010 SN:000200016870 | P0; 450nm; PCB
2010-10-31 | 0:00 | 0.654 | 22.2 | PQ2 | Joel Guerrero | DR/2010 SN:000200016870 | P0; 450nm; PCB
2010-10-31 | 0:00 | 0.754 | 17.6 | PQ3 | Joel Guerrero | DR/2010 SN:000200016870 | P0; 450nm; PCB
2010-10-31 | 0:00 | 0.779 | 16.6 | DD6-7 | Joel Guerrero | DR/2010 SN:000200016870 | P0; 450nm; PCB
2010-10-31 | 0:00 | 0.852 | 14 | DD6-1/4 COMP | Joel Guerrero | DR/2010 SN:000200016870 | P0; 450nm; PCB
2010-10-31 | 0:00 | 0.852 | 14.1 | P7-T1-5 | Joel Guerrero | DR/2010 SN:000200016870 | P0; 450nm; PCB
2010-10-31 | 0:00 | 1.007 | 9.84 | P7-T3-5 | Joel Guerrero | DR/2010 SN:000200016870 | P0; 450nm; PCB
2010-10-31 | 0:00 | 0.624 | 23.7 | P7-T5-2 | Joel Guerrero | DR/2010 SN:000200016870 | P0; 450nm; PCB
2010-10-31 | 0:00 | 0.237 | 58 | P7-T6-3 | Joel Guerrero | DR/2010 SN:000200016870 | P0; 450nm; PCB
2010-10-31 | 0:00 | 0.911 | 12.3 | P7-T9-1 | Joel Guerrero | DR/2010 SN:000200016870 | P0; 450nm; PCB
2010-10-31 | 0:00 | 0.92 | 12 | P7-T10-6 | Joel Guerrero | DR/2010 SN:000200016870 | P0; 450nm; PCB
2010-10-31 | 0:00 | 0.126 | 74.8 | NIST1939a | Joel Guerrero | DR/2010 SN:000200016870 | P0; 450nm; PCB
2010-10-31 | 0:00 | 0.789 | 16.3 | HS1 | Joel Guerrero | DR/2010 SN:000200016870 | P0; 450nm; PCB

## QAQC
Quality Control: Control standards in methanol with certified values of 3 ± 0.6 ppb PCBs (as Aroclor 1254)
are provided with each ImmunoAssay kit. These standard reference materials (SRMs) are included in every batch run of 40-60 test tubes.
Each SRM is treated in the same manner as an unknown sample.
Calib. Curve # | PCB (Aroclor 1254) | Analyte (mg/Kg) | Mean | Certified | % Recov | % RPD
1 | 3.1539 | PCB | 3.1537 | 3 | 105.1% | 5.123%
2 | 3.0607
3 | 3.1117
4 | 3.2884
5
6
7
8
9
10
11
12
13
14
15
16
Mean | 3.1537
Stdev | 0.097574
%RSD | 3.094%
Definitions:
Mean: Arithmethic Mean
Stdev: Standard Deviation from duplicate assay analyses (n=2...4)
% RSD: Relative Standard Deviation whereby; [{stdev/mean} * 100]
% Recov: Percent recovery of assay result from certified value [{IA result / cert} * 100]
% RPD: = Relative Percent Difference whereby; [abs{cert-mean}/cert * 100]

## SDI RaPID Assay Kit Lot#
| Lot # | Exp. Date | Date Opened
PCB RaPID Assay Kit: |  |  | 2010-10-31
Components: |  |  | 2010-10-31
Washing Solution |  |  | 2010-10-31
Stopping Solution |  |  | 2010-10-31
Color Solution |  |  | 2010-10-31
Magnetic Particles |  |  | 2010-10-31
Enzyme Conjugate |  |  | 2010-10-31
Diluent/0.0 ppb Standard |  |  | 2010-10-31
Aroclor 1254 Std 1 (0.25 ppb) |  |  | 2010-10-31
Aroclor 1254 Std 2 (1.0 ppb) |  |  | 2010-10-31
Aroclor 1254 Std 3 (5.0 ppb) |  |  | 2010-10-31
Control (3.0 ppb) |  |  | 2010-10-31
Sample Diluent |  |  | 2010-10-31

## PCB Calib Curve #1 (31Oct10)
Date: | 2010-10-31
ImmunoAssay (IA) Method Calibration for PCBs Using a Direct-Reading (DR) Spectrophotometer [HACH]
conc | abs | (r/r0)/(1-r/r0) | ln Conc | LgtB
0 | 1.166
0.25 | 0.902 | 3.4167 | -1.3863 | 1.2287
1 | 0.664 | 1.3227 | 0 | 0.27968
5 | 0.467 | 0.6681 | 1.6094 | -0.40332
x = ln (conc)
| b = | 0.4086
| m = | -0.5416
x = (y -b) / m
y = ln (r/r0)/(1-r/r0)
x = ((ln((r/r0)/(1-(r/r0)))-b))/m
conc = exp(x) = exp(((ln((r/r0)/(1-r/r0))-b))/m)
LABEL ID | abs | x | conc | % recov | pcb
control [aroclor 1254] | 0.521 | 1.1486 | 3.1539 | 105.1%
RB | 1.075 | -3.8047 | 0.022266 |  | 4.4978
P7-T1-1 | 0.78 | -0.54442 | 0.58018 |  | 234.3457
P7-T1-2 | 0.707 | -0.043169 | 0.95775 |  | 191.9109
P7-T1-3 | 0.814 | -0.79344 | 0.45228 |  | 90.8979
P7-T1-4 | 0.898 | -1.4782 | 0.22805 |  | 45.6964
P7-T1-5 | 0.777 | -0.52301 | 0.59274 |  | 118.8648
P7-T2-1 | 0.72 | -0.12986 | 0.87822 |  | 176.2019
P7-T2-2 | 0.743 | -0.28568 | 0.7515 |  | 150.7784
P7-T2-3 | 0.95 | -1.9804 | 0.13801 |  | 27.7236
P7-T2-4 | 0.661 | 0.25739 | 1.2936 |  | 260.9067
P7-T2-5 | 0.871 | -1.2446 | 0.28806 |  | 58.049
P7-T3-1 | 0.726 | -0.17019 | 0.8435 |  | 169.7766
P7-T3-2 | 0.806 | -0.73371 | 0.48012 |  | 96.2914
P7-T3-3 | 0.647 | 0.34741 | 1.4154 |  | 284.7996
P7-T3-4 | 0.84 | -0.99318 | 0.3704 |  | 74.2266
P7-T3-5 | 0.94 | -1.8773 | 0.153 |  | 30.7375
P7-T4-1 | 0.619 | 0.52611 | 1.6923 |  | 340.6951
P7-T4-2 | 0.754 | -0.36146 | 0.69666 |  | 140.4015
P7-T4-3 | 0.784 | -0.5731 | 0.56378 |  | 113.1814
P7-T4-4 | 0.762 | -0.41716 | 0.65892 |  | 133.1015
P7-T4-5 | 1.025 | -2.9082 | 0.054573 |  | 10.9711
Note:
abs: measured absorbance (at l=450 nm)
% recov: percent recovery of assay result from certified value [{IA result / cert} * 100]
conc: analyte concentration (mg/Kg) (ppb) from ImmunoAssay (IA)
pcb: dry-weight & dilution corrected analyte concentration (mg/Kg) [{VMeOH * conc * df } / drywt]
label id: P7 = Pier 7; Tx-x = Tansect № (1…10) - ID № (1….6)

## PCB Calib Curve #2 (31Oct10)
Date: | 2010-10-31
ImmunoAssay (IA) Method Calibration for PCBs Using a Direct-Reading (DR) Spectrophotometer [HACH]
conc | abs | (r/r0)/(1-r/r0) | ln Conc | LgtB
0 | 1.06
0.25 | 0.866 | 4.4639 | -1.3863 | 1.496
1 | 0.643 | 1.542 | 0 | 0.43306
5 | 0.451 | 0.74056 | 1.6094 | -0.30035
x = ln (conc)
| b = | 0.5872
| m = | -0.5958
x = (y -b) / m
y = ln (r/r0)/(1-r/r0)
x = ((ln((r/r0)/(1-(r/r0)))-b))/m
conc = exp(x) = exp(((ln((r/r0)/(1-r/r0))-b))/m)
LABEL ID | abs | x | conc | % recov | pcb
control [aroclor 1254] | 0.509 | 1.1186 | 3.0607 | 102%
MB | 1.034 | -5.1962 | 0.0055376 |  | 1.105
P7-T5-1 | 0.763 | -0.59806 | 0.54988 |  | 111.0749
P7-T5-2 | 0.529 | 0.9919 | 2.6964 |  | 539.9652
P7-T5-3 | 0.715 | -0.23756 | 0.78855 |  | 159.287
P7-T5-4 | 0.723 | -0.29561 | 0.74408 |  | 150.0182
P7-T5-5 | 0.812 | -1.0052 | 0.36599 |  | 73.7231
P7-T6-1 | 0.653 | 0.19207 | 1.2118 |  | 243.2908
P7-T6-2 | 0.643 | 0.25871 | 1.2953 |  | 261.6431
P7-T6-3 | 0.21 | 3.3322 | 28.0001 |  | 5643.5971
P7-T6-4 | 0.619 | 0.41648 | 1.5166 |  | 305.3487
P7-T6-5 | 0.562 | 0.78264 | 2.1872 |  | 438.6649
P7-T7-1 | 0.666 | 0.1045 | 1.1102 |  | 224.2514
P7-T7-2 | 0.643 | 0.25871 | 1.2953 |  | 261.4078
P7-T7-3 | 0.743 | -0.4441 | 0.6414 |  | 128.9571
P7-T7-4 | 0.5 | 1.1758 | 3.2407 |  | 654.6141
P7-T7-5 | 0.688 | -0.046484 | 0.95458 |  | 192.8252
P7-T7-6 | 0.769 | -0.64546 | 0.52442 |  | 104.9565
P7-T8-1 | 0.664 | 0.11805 | 1.1253 |  | 226.743
P7-T8-2 | 0.743 | -0.4441 | 0.6414 |  | 129.253
P7-T8-3 | 0.711 | -0.2088 | 0.81156 |  | 162.6665
HS1 | 0.571 | 0.72537 | 2.0655 |  | 414.7401
Note:
abs: measured absorbance (at l=450 nm)
% recov: percent recovery of assay result from certified value [{IA result / cert} * 100]
conc: analyte concentration (mg/Kg) (ppb) from ImmunoAssay (IA)
pcb: dry-weight & dilution corrected analyte concentration (mg/Kg) [{VMeOH * conc * df } / drywt]
label id: P7 = Pier 7; Tx-x = Tansect № (1…10) - ID № (1….6)

## PCB Calib Curve #3 (31Oct10)
Date: | 2010-10-31
ImmunoAssay (IA) Method Calibration for PCBs Using a Direct-Reading (DR) Spectrophotometer [HACH]
conc | abs | (r/r0)/(1-r/r0) | ln Conc | LgtB
0 | 1.438
0.25 | 1.096 | 3.2047 | -1.3863 | 1.1646
1 | 0.784 | 1.1988 | 0 | 0.1813
5 | 0.531 | 0.58545 | 1.6094 | -0.53538
x = ln (conc)
| b = | 0.3121
| m = | -0.5642
x = (y -b) / m
y = ln (r/r0)/(1-r/r0)
x = ((ln((r/r0)/(1-(r/r0)))-b))/m
conc = exp(x) = exp(((ln((r/r0)/(1-r/r0))-b))/m)
LABEL ID | abs | x | conc | % recov | pcb
control [aroclor 1254] | 0.602 | 1.1352 | 3.1117 | 103.7%
RB | 1.304 | -3.4797 | 0.030816 |  | 6.2057
P7-T9-1 | 1.067 | -1.3192 | 0.26734 |  | 53.7829
P7-T9-2 | 1.016 | -1.0041 | 0.36637 |  | 73.888
P7-T9-3 | 1.015 | -0.99818 | 0.36855 |  | 74.3209
P7-T9-4 | 1.017 | -1.0101 | 0.36419 |  | 73.3692
P7-T9-5 | 0.993 | -0.86947 | 0.41917 |  | 84.3271
P7-T9-6 | 1.001 | -0.91585 | 0.40018 |  | 80.1227
P7-T10-1 | 0.874 | -0.22319 | 0.79996 |  | 160.8839
P7-T10-2 | 1.18 | -2.1415 | 0.11748 |  | 23.5901
P7-T10-3 | 0.92 | -0.4649 | 0.6282 |  | 126.4404
P7-T10-4 | 0.978 | -0.78373 | 0.4567 |  | 91.5388
P7-T10-5 | 0.908 | -0.40104 | 0.66962 |  | 134.3635
P7-T10-6 | 0.947 | -0.61105 | 0.54278 |  | 109.1287
Note:
abs: measured absorbance (at l=450 nm)
% recov: percent recovery of assay result from certified value [{IA result / cert} * 100]
conc: analyte concentration (mg/Kg) (ppb) from ImmunoAssay (IA)
pcb: dry-weight & dilution corrected analyte concentration (mg/Kg) [{VMeOH * conc * df } / drywt]
label id: P7 = Pier 7; Tx-x = Tansect № (1…10) - ID № (1….6)

## PCB Calib Curve #4 (31Oct10)
Date: | 2010-10-31
ImmunoAssay (IA) Method Calibration for PCBs Using a Direct-Reading (DR) Spectrophotometer [HACH]
conc | abs | (r/r0)/(1-r/r0) | ln Conc | LgtB
0 | 1.297
0.25 | 1.043 | 4.1063 | -1.3863 | 1.4125
1 | 0.833 | 1.7953 | 0 | 0.58515
5 | 0.546 | 0.72703 | 1.6094 | -0.31879
x = ln (conc)
| b = | 0.6026
| m = | -0.5775
x = (y -b) / m
y = ln (r/r0)/(1-r/r0)
x = ((ln((r/r0)/(1-(r/r0)))-b))/m
conc = exp(x) = exp(((ln((r/r0)/(1-r/r0))-b))/m)
LABEL ID | abs | x | conc | % recov | pcb | pah corr | pcb corr
control [aroclor 1254] | 0.621 | 1.1904 | 3.2884 | 109.6%
MB | 1.288 | -7.5516 | 0.00052529 |  | 0.10482
PQ1 | 0.75 | 0.49693 | 1.6437 |  | 330.2367
PQ2 | 0.654 | 1.0141 | 2.7569 |  | 549.8468
PQ3 | 0.754 | 0.47501 | 1.608 |  | 394.6797
DD6-7 | 0.779 | 0.33691 | 1.4006 |  | 282.4712
DD6-1/4 | 0.852 | -0.081233 | 0.92198 |  | 185.4977
P7-T1-5 | 0.852 | -0.081233 | 0.92198 |  | 184.89
P7-T3-5 | 1.007 | -1.1121 | 0.32886 |  | 66.0665
P7-T5-2 | 0.624 | 1.1744 | 3.2361 |  | 648.0507
P7-T6-3 | 0.237 | 3.6373 | 37.9907 |  | 7657.2673
P7-T9-1 | 0.911 | -0.44347 | 0.6418 |  | 129.115
P7-T10-6 | 0.92 | -0.50135 | 0.60571 |  | 121.7818
NIST1939a | 0.126 | 4.9038 | 134.798 |  | 107677.2925
HS1 | 0.789 | 0.28107 | 1.3245 |  | 265.9613
Note:
abs: measured absorbance (at l=450 nm)
% recov: percent recovery of assay result from certified value [{IA result / cert} * 100]
conc: analyte concentration (mg/Kg) (ppb) from ImmunoAssay (IA)
pcb: dry-weight & dilution corrected analyte concentration (mg/Kg) [{VMeOH * conc * df } / drywt]
label id: P7 = Pier 7; Tx-x = Tansect № (1…10) - ID № (1….6)

## PCB Assay Meta Data
SAMPLE
LABEL | Dry wt | VMeOH | Vext | Vdil | df
RB | 10 | 20 | 0.05 | 5 | 101
P7-T1-1 | 5.001 | 20 | 0.05 | 5 | 101
P7-T1-2 | 10.081 | 20 | 0.05 | 5 | 101
P7-T1-3 | 10.051 | 20 | 0.05 | 5 | 101
P7-T1-4 | 10.081 | 20 | 0.05 | 5 | 101
P7-T1-5 | 10.073 | 20 | 0.05 | 5 | 101
P7-T2-1 | 10.068 | 20 | 0.05 | 5 | 101
P7-T2-2 | 10.068 | 20 | 0.05 | 5 | 101
P7-T2-3 | 10.056 | 20 | 0.05 | 5 | 101
P7-T2-4 | 10.015 | 20 | 0.05 | 5 | 101
P7-T2-5 | 10.024 | 20 | 0.05 | 5 | 101
P7-T3-1 | 10.036 | 20 | 0.05 | 5 | 101
P7-T3-2 | 10.072 | 20 | 0.05 | 5 | 101
P7-T3-3 | 10.039 | 20 | 0.05 | 5 | 101
P7-T3-4 | 10.08 | 20 | 0.05 | 5 | 101
P7-T3-5 | 10.055 | 20 | 0.05 | 5 | 101
P7-T4-1 | 10.034 | 20 | 0.05 | 5 | 101
P7-T4-2 | 10.023 | 20 | 0.05 | 5 | 101
P7-T4-3 | 10.062 | 20 | 0.05 | 5 | 101
P7-T4-4 | 10 | 20 | 0.05 | 5 | 101
P7-T4-5 | 10.048 | 20 | 0.05 | 5 | 101
MB | 10.123 | 20 | 0.05 | 5 | 101
P7-T5-1 | 10 | 20 | 0.05 | 5 | 101
P7-T5-2 | 10.087 | 20 | 0.05 | 5 | 101
P7-T5-3 | 10 | 20 | 0.05 | 5 | 101
P7-T5-4 | 10.019 | 20 | 0.05 | 5 | 101
P7-T5-5 | 10.028 | 20 | 0.05 | 5 | 101
P7-T6-1 | 10.061 | 20 | 0.05 | 5 | 101
P7-T6-2 | 10 | 20 | 0.05 | 5 | 101
P7-T6-3 | 10.022 | 20 | 0.05 | 5 | 101
P7-T6-4 | 10.033 | 20 | 0.05 | 5 | 101
P7-T6-5 | 10.072 | 20 | 0.05 | 5 | 101
P7-T7-1 | 10 | 20 | 0.05 | 5 | 101
P7-T7-2 | 10.009 | 20 | 0.05 | 5 | 101
P7-T7-3 | 10.047 | 20 | 0.05 | 5 | 101
P7-T7-4 | 10 | 20 | 0.05 | 5 | 101
P7-T7-5 | 10 | 20 | 0.05 | 5 | 101
P7-T7-6 | 10.093 | 20 | 0.05 | 5 | 101
P7-T8-1 | 10.025 | 20 | 0.05 | 5 | 101
P7-T8-2 | 10.024 | 20 | 0.05 | 5 | 101
P7-T8-3 | 10.078 | 20 | 0.05 | 5 | 101
HS1 | 10.06 | 20 | 0.05 | 5 | 101
Ave: | 9.9226
VMeOH: Volume of extractant (methanol) [ml]
Vext: Volume of dilution extract (ml)
Vdil: Volume of diluent (ml)
df: Dilution factor [(Vext + Vdil)/Vext]
Label: P7 = Pier 7; Tx-x = Tansect № (1…10) - ID № (1….6)
SAMPLE
LABEL | Dry wt | VMeOH | Vext | Vdil | df
RB | 10.031 | 20 | 0.05 | 5 | 101
P7-T9-1 | 10.041 | 20 | 0.05 | 5 | 101
P7-T9-2 | 10.016 | 20 | 0.05 | 5 | 101
P7-T9-3 | 10.017 | 20 | 0.05 | 5 | 101
P7-T9-4 | 10.027 | 20 | 0.05 | 5 | 101
P7-T9-5 | 10.041 | 20 | 0.05 | 5 | 101
P7-T9-6 | 10.089 | 20 | 0.05 | 5 | 101
P7-T10-1 | 10.044 | 20 | 0.05 | 5 | 101
P7-T10-2 | 10.06 | 20 | 0.05 | 5 | 101
P7-T10-3 | 10.036 | 20 | 0.05 | 5 | 101
P7-T10-4 | 10.078 | 20 | 0.05 | 5 | 101
P7-T10-5 | 10.067 | 20 | 0.05 | 5 | 101
P7-T10-6 | 10.047 | 20 | 0.05 | 5 | 101
MB | 10.123 | 20 | 0.05 | 5 | 101
PQ1 | 5.027 | 10 | 0.05 | 5 | 101
PQ2 | 5.064 | 10 | 0.05 | 5 | 101
PQ3 | 4.115 | 10 | 0.05 | 5 | 101
DD6-7 | 5.008 | 10 | 0.05 | 5 | 101
DD6-1/4 | 5.02 | 10 | 0.05 | 5 | 101
P7-T1-5 | 10.073 | 20 | 0.05 | 5 | 101
P7-T3-5 | 10.055 | 20 | 0.05 | 5 | 101
P7-T5-2 | 10.087 | 20 | 0.05 | 5 | 101
P7-T6-3 | 10.022 | 20 | 0.05 | 5 | 101
P7-T9-1 | 10.041 | 20 | 0.05 | 5 | 101
P7-T10-6 | 10.047 | 20 | 0.05 | 5 | 101
NIST1939a | 10.04 | 20 | 0.05 | 20 | 401
HS1 | 10.06 | 20 | 0.05 | 5 | 101
Ave: | 9.088
VMeOH: Volume of extractant (methanol) [ml]
Vext: Volume of dilution extract (ml)
Vdil: Volume of diluent (ml)
df: Dilution factor [(Vext + Vdil)/Vext]
Label: P7 = Pier 7; Tx-x = Tansect № (1…10) - Sample № (1….6)

## 31Oct10
DATE | TIME | ABS | %T | SAMPLE | OPERATOR | INSTRUMENT | NOTES
2010-10-31 | 0:00 | 0 | 100 | zero | Joel Guerrero | DR/2010 SN:000200016870 | P0; 450nm; PCB
2010-10-31 | 0:00 | 1.166 | 6.83 | 0.0 std | Joel Guerrero | DR/2010 SN:000200016870 | P0; 450nm; PCB
2010-10-31 | 0:00 | 0.902 | 12.5 | 0.25 std | Joel Guerrero | DR/2010 SN:000200016870 | P0; 450nm; PCB
2010-10-31 | 0:00 | 0.664 | 21.7 | 1.00 std | Joel Guerrero | DR/2010 SN:000200016870 | P0; 450nm; PCB
2010-10-31 | 0:00 | 0.467 | 34.1 | 5.00 std | Joel Guerrero | DR/2010 SN:000200016870 | P0; 450nm; PCB
2010-10-31 | 0:00 | 0.521 | 30.1 | control | Joel Guerrero | DR/2010 SN:000200016870 | P0; 450nm; PCB
2010-10-31 | 0:00 | 1.075 | 8.42 | RB | Joel Guerrero | DR/2010 SN:000200016870 | P0; 450nm; PCB
2010-10-31 | 0:00 | 0.78 | 16.6 | P7-T1-1 | Joel Guerrero | DR/2010 SN:000200016870 | P0; 450nm; PCB
2010-10-31 | 0:00 | 0.707 | 19.6 | P7-T1-2 | Joel Guerrero | DR/2010 SN:000200016870 | P0; 450nm; PCB
2010-10-31 | 0:00 | 0.814 | 15.4 | P7-T1-3 | Joel Guerrero | DR/2010 SN:000200016870 | P0; 450nm; PCB
2010-10-31 | 0:00 | 0.898 | 12.6 | P7-T1-4 | Joel Guerrero | DR/2010 SN:000200016870 | P0; 450nm; PCB
2010-10-31 | 0:00 | 0.777 | 16.7 | P7-T1-5 | Joel Guerrero | DR/2010 SN:000200016870 | P0; 450nm; PCB
2010-10-31 | 0:00 | 0.72 | 19 | P7-T2-1 | Joel Guerrero | DR/2010 SN:000200016870 | P0; 450nm; PCB
2010-10-31 | 0:00 | 0.743 | 18.1 | P7-T2-2 | Joel Guerrero | DR/2010 SN:000200016870 | P0; 450nm; PCB
2010-10-31 | 0:00 | 0.95 | 11.2 | P7-T2-3 | Joel Guerrero | DR/2010 SN:000200016870 | P0; 450nm; PCB
2010-10-31 | 0:00 | 0.661 | 21.8 | P7-T2-4 | Joel Guerrero | DR/2010 SN:000200016870 | P0; 450nm; PCB
2010-10-31 | 0:00 | 0.871 | 13.4 | P7-T2-5 | Joel Guerrero | DR/2010 SN:000200016870 | P0; 450nm; PCB
2010-10-31 | 0:00 | 0.726 | 18.8 | P7-T3-1 | Joel Guerrero | DR/2010 SN:000200016870 | P0; 450nm; PCB
2010-10-31 | 0:00 | 0.806 | 15.6 | P7-T3-2 | Joel Guerrero | DR/2010 SN:000200016870 | P0; 450nm; PCB
2010-10-31 | 0:00 | 0.647 | 22.5 | P7-T3-3 | Joel Guerrero | DR/2010 SN:000200016870 | P0; 450nm; PCB
2010-10-31 | 0:00 | 0.84 | 14.4 | P7-T3-4 | Joel Guerrero | DR/2010 SN:000200016870 | P0; 450nm; PCB
2010-10-31 | 0:00 | 0.94 | 11.5 | P7-T3-5 | Joel Guerrero | DR/2010 SN:000200016870 | P0; 450nm; PCB
2010-10-31 | 0:00 | 0.619 | 24 | P7-T4-1 | Joel Guerrero | DR/2010 SN:000200016870 | P0; 450nm; PCB
2010-10-31 | 0:00 | 0.754 | 17.6 | P7-T4-2 | Joel Guerrero | DR/2010 SN:000200016870 | P0; 450nm; PCB
2010-10-31 | 0:00 | 0.784 | 16.4 | P7-T4-3 | Joel Guerrero | DR/2010 SN:000200016870 | P0; 450nm; PCB
2010-10-31 | 0:00 | 0.762 | 17.3 | P7-T4-4 | Joel Guerrero | DR/2010 SN:000200016870 | P0; 450nm; PCB
2010-10-31 | 0:00 | 1.025 | 9.45 | P7-T4-5 | Joel Guerrero | DR/2010 SN:000200016870 | P0; 450nm; PCB
2010-10-31 | 0:00 | 1.06 | 8.7 | 0.00 std | Joel Guerrero | DR/2010 SN:000200016870 | P0; 450nm; PCB
2010-10-31 | 0:00 | 0.866 | 13.6 | 0.25 std | Joel Guerrero | DR/2010 SN:000200016870 | P0; 450nm; PCB
2010-10-31 | 0:00 | 0.643 | 22.8 | 1.00 std | Joel Guerrero | DR/2010 SN:000200016870 | P0; 450nm; PCB
2010-10-31 | 0:00 | 0.451 | 35.4 | 5.00 std | Joel Guerrero | DR/2010 SN:000200016870 | P0; 450nm; PCB
2010-10-31 | 0:00 | 0.509 | 31 | control | Joel Guerrero | DR/2010 SN:000200016870 | P0; 450nm; PCB
2010-10-31 | 0:00 | 1.034 | 9.25 | MB | Joel Guerrero | DR/2010 SN:000200016870 | P0; 450nm; PCB
2010-10-31 | 0:00 | 0.763 | 17.3 | P7-T5-1 | Joel Guerrero | DR/2010 SN:000200016870 | P0; 450nm; PCB
2010-10-31 | 0:00 | 0.529 | 29.6 | P7-T5-2 | Joel Guerrero | DR/2010 SN:000200016870 | P0; 450nm; PCB
2010-10-31 | 0:00 | 0.715 | 19.3 | P7-T5-3 | Joel Guerrero | DR/2010 SN:000200016870 | P0; 450nm; PCB
2010-10-31 | 0:00 | 0.723 | 18.9 | P7-T5-4 | Joel Guerrero | DR/2010 SN:000200016870 | P0; 450nm; PCB
2010-10-31 | 0:00 | 0.812 | 15.4 | P7-T5-5 | Joel Guerrero | DR/2010 SN:000200016870 | P0; 450nm; PCB
2010-10-31 | 0:00 | 0.653 | 22.2 | P7-T6-1 | Joel Guerrero | DR/2010 SN:000200016870 | P0; 450nm; PCB
2010-10-31 | 0:00 | 0.643 | 22.7 | P7-T6-2 | Joel Guerrero | DR/2010 SN:000200016870 | P0; 450nm; PCB
2010-10-31 | 0:00 | 0.21 | 61.6 | P7-T6-3 | Joel Guerrero | DR/2010 SN:000200016870 | P0; 450nm; PCB
2010-10-31 | 0:00 | 0.619 | 24 | P7-T6-4 | Joel Guerrero | DR/2010 SN:000200016870 | P0; 450nm; PCB
2010-10-31 | 0:00 | 0.562 | 27.4 | P7-T6-5 | Joel Guerrero | DR/2010 SN:000200016870 | P0; 450nm; PCB
2010-10-31 | 0:00 | 0.666 | 21.6 | P7-T7-1 | Joel Guerrero | DR/2010 SN:000200016870 | P0; 450nm; PCB
2010-10-31 | 0:00 | 0.643 | 22.7 | P7-T7-2 | Joel Guerrero | DR/2010 SN:000200016870 | P0; 450nm; PCB
2010-10-31 | 0:00 | 0.743 | 18.1 | P7-T7-3 | Joel Guerrero | DR/2010 SN:000200016870 | P0; 450nm; PCB
2010-10-31 | 0:00 | 0.5 | 31.6 | P7-T7-4 | Joel Guerrero | DR/2010 SN:000200016870 | P0; 450nm; PCB
2010-10-31 | 0:00 | 0.688 | 20.5 | P7-T7-5 | Joel Guerrero | DR/2010 SN:000200016870 | P0; 450nm; PCB
2010-10-31 | 0:00 | 0.769 | 17 | P7-T7-6 | Joel Guerrero | DR/2010 SN:000200016870 | P0; 450nm; PCB
2010-10-31 | 0:00 | 0.664 | 21.7 | P7-T8-1 | Joel Guerrero | DR/2010 SN:000200016870 | P0; 450nm; PCB
2010-10-31 | 0:00 | 0.743 | 18.1 | P7-T8-2 | Joel Guerrero | DR/2010 SN:000200016870 | P0; 450nm; PCB
2010-10-31 | 0:00 | 0.711 | 19.4 | P7-T8-3 | Joel Guerrero | DR/2010 SN:000200016870 | P0; 450nm; PCB
2010-10-31 | 0:00 | 0.571 | 26.9 | HS1 | Joel Guerrero | DR/2010 SN:000200016870 | P0; 450nm; PCB
2010-10-31 | 0:00 | 0 | 100 | zero | Joel Guerrero | DR/2010 SN:000200016870 | P0; 450nm; PCB
2010-10-31 | 0:00 | 1.438 | 3.65 | 0.00 std | Joel Guerrero | DR/2010 SN:000200016870 | P0; 450nm; PCB
2010-10-31 | 0:00 | 1.096 | 8.03 | 0.25 std | Joel Guerrero | DR/2010 SN:000200016870 | P0; 450nm; PCB
2010-10-31 | 0:00 | 0.784 | 16.5 | 1.00 std | Joel Guerrero | DR/2010 SN:000200016870 | P0; 450nm; PCB
2010-10-31 | 0:00 | 0.531 | 29.5 | 5.00 std | Joel Guerrero | DR/2010 SN:000200016870 | P0; 450nm; PCB
2010-10-31 | 0:00 | 0.602 | 25 | control | Joel Guerrero | DR/2010 SN:000200016870 | P0; 450nm; PCB
2010-10-31 | 0:00 | 1.304 | 4.96 | RB | Joel Guerrero | DR/2010 SN:000200016870 | P0; 450nm; PCB
2010-10-31 | 0:00 | 1.067 | 8.56 | P7-T9-1 | Joel Guerrero | DR/2010 SN:000200016870 | P0; 450nm; PCB
2010-10-31 | 0:00 | 1.016 | 9.65 | P7-T9-2 | Joel Guerrero | DR/2010 SN:000200016870 | P0; 450nm; PCB
2010-10-31 | 0:00 | 1.015 | 9.66 | P7-T9-3 | Joel Guerrero | DR/2010 SN:000200016870 | P0; 450nm; PCB
2010-10-31 | 0:00 | 1.017 | 9.61 | P7-T9-4 | Joel Guerrero | DR/2010 SN:000200016870 | P0; 450nm; PCB
2010-10-31 | 0:00 | 0.993 | 10.2 | P7-T9-5 | Joel Guerrero | DR/2010 SN:000200016870 | P0; 450nm; PCB
2010-10-31 | 0:00 | 1.001 | 9.99 | P7-T9-6 | Joel Guerrero | DR/2010 SN:000200016870 | P0; 450nm; PCB
2010-10-31 | 0:00 | 0.874 | 13.4 | P7-T10-1 | Joel Guerrero | DR/2010 SN:000200016870 | P0; 450nm; PCB
2010-10-31 | 0:00 | 1.18 | 6.61 | P7-T10-2 | Joel Guerrero | DR/2010 SN:000200016870 | P0; 450nm; PCB
2010-10-31 | 0:00 | 0.92 | 12 | P7-T10-3 | Joel Guerrero | DR/2010 SN:000200016870 | P0; 450nm; PCB
2010-10-31 | 0:00 | 0.978 | 10.5 | P7-T10-4 | Joel Guerrero | DR/2010 SN:000200016870 | P0; 450nm; PCB
2010-10-31 | 0:00 | 0.908 | 12.4 | P7-T10-5 | Joel Guerrero | DR/2010 SN:000200016870 | P0; 450nm; PCB
2010-10-31 | 0:00 | 0.947 | 11.3 | P7-T10-6 | Joel Guerrero | DR/2010 SN:000200016870 | P0; 450nm; PCB
2010-10-31 | 0:00 | 1.297 | 5.05 | 0.00 std | Joel Guerrero | DR/2010 SN:000200016870 | P0; 450nm; PCB
2010-10-31 | 0:00 | 1.043 | 9.06 | 0.25 std | Joel Guerrero | DR/2010 SN:000200016870 | P0; 450nm; PCB
2010-10-31 | 0:00 | 0.833 | 14.7 | 1.00 std | Joel Guerrero | DR/2010 SN:000200016870 | P0; 450nm; PCB
2010-10-31 | 0:00 | 0.546 | 28.5 | 5.00 std | Joel Guerrero | DR/2010 SN:000200016870 | P0; 450nm; PCB
2010-10-31 | 0:00 | 0.621 | 24 | control | Joel Guerrero | DR/2010 SN:000200016870 | P0; 450nm; PCB
2010-10-31 | 0:00 | 1.288 | 5.07 | MB | Joel Guerrero | DR/2010 SN:000200016870 | P0; 450nm; PCB
2010-10-31 | 0:00 | 0.75 | 17.8 | PQ1 | Joel Guerrero | DR/2010 SN:000200016870 | P0; 450nm; PCB
2010-10-31 | 0:00 | 0.654 | 22.2 | PQ2 | Joel Guerrero | DR/2010 SN:000200016870 | P0; 450nm; PCB
2010-10-31 | 0:00 | 0.754 | 17.6 | PQ3 | Joel Guerrero | DR/2010 SN:000200016870 | P0; 450nm; PCB
2010-10-31 | 0:00 | 0.779 | 16.6 | DD6-7 | Joel Guerrero | DR/2010 SN:000200016870 | P0; 450nm; PCB
2010-10-31 | 0:00 | 0.852 | 14 | DD6-1/4 COMP | Joel Guerrero | DR/2010 SN:000200016870 | P0; 450nm; PCB
2010-10-31 | 0:00 | 0.852 | 14.1 | P7-T1-5 | Joel Guerrero | DR/2010 SN:000200016870 | P0; 450nm; PCB
2010-10-31 | 0:00 | 1.007 | 9.84 | P7-T3-5 | Joel Guerrero | DR/2010 SN:000200016870 | P0; 450nm; PCB
2010-10-31 | 0:00 | 0.624 | 23.7 | P7-T5-2 | Joel Guerrero | DR/2010 SN:000200016870 | P0; 450nm; PCB
2010-10-31 | 0:00 | 0.237 | 58 | P7-T6-3 | Joel Guerrero | DR/2010 SN:000200016870 | P0; 450nm; PCB
2010-10-31 | 0:00 | 0.911 | 12.3 | P7-T9-1 | Joel Guerrero | DR/2010 SN:000200016870 | P0; 450nm; PCB
2010-10-31 | 0:00 | 0.92 | 12 | P7-T10-6 | Joel Guerrero | DR/2010 SN:000200016870 | P0; 450nm; PCB
2010-10-31 | 0:00 | 0.126 | 74.8 | NIST1939a | Joel Guerrero | DR/2010 SN:000200016870 | P0; 450nm; PCB
2010-10-31 | 0:00 | 0.789 | 16.3 | HS1 | Joel Guerrero | DR/2010 SN:000200016870 | P0; 450nm; PCB